KTIA_NAP_13-1-2013-0001: „Agykutatási kiválósági központok fejlesztése”

A Szegedi Tudományegyetem Fluoreszcens bővítés AxioImager M1 motoros mikroszkóphoz eszköz/szolgáltatás/fogyóeszköz beszerzés műszaki minimumfeltételei:
	Megnevezés
	Specifikáció
	Mennyiség

	Fluoreszcens bővítés AxioImager M1 motoros mikroszkóphoz:
Meglévő Zeiss AxioImager M1 motoros fénymikroszkóp fluoreszcens bővítése Rhodamin, CY3, DAPI és FITC csatornákra

Gyártó: Carl Zeiss Technika Kft.

	6-részes motoros reflektorrevolver P&C modulokhoz

FL és HD ráesőfényű megvilágítás elektronikus fényzárral (shutter) és manuális mezőblendével. Áteresztési képesség: 300-1000 nm. 300 nm -nél > 40%, és 1000 nm-nél > 60%

Nagy sebességű fényzár,"Axio Imager" M és Z mikroszkópokhoz. Kapcsolási idő < 20 ms,kapcsolási frekvencia max. 5 Hz. Firmware Version 06.008 vagy magasabb

Centrálható mezőblende, állítható stop pozícióval

FL attenuator, diszkrét intenzitás csökkentő kerék. Szűrőkerék 6 pozícióval: Pos. 1: sima apertúra.

Pos. 2-6: attenuation 70, 50, 40, 20 és 2% áteresztéssel

Reflektor modul FL EC P&C reflexiomentes kiképzéssel az emissziós oldalon max. szűrővastagság 5 mm

Szűrőkészlet 49 DAPI eltolódásmentes, EX G 365, BS FT 395, EM BP 445/50

Szűrőkészlet 38 Endow GFP, FITC eltolódásmentes, EX BP 470/40, BS FT 495, EM BP 525/50

Szűrőkészlet 20 Rhodamin eltolódásmentes, EX BP 546/12, BS FT 560, EM BP 575-640

Szűrőkészlet 43 Cy 3, d=25 eltolódás mentes, EX BP 545/25, BS FT 570, EM BP 605/70

Kompenzátor befogó 6x20

Fluoreszcens védőernyő

HBO 100-as lámpaház, lámpafoglalattal és kollektorral, beépített hőszűrővel és 2m csatlakozó kábellel

HBO 103W nagynyomású higanygőzlámpa

Tápegység HBO 100, 90...250V, 50...60Hz, 265 VA, nullázható számlálóval az előlapon

Beállítátó optika HBO/XBO "Axio Imager"-hez

Kamera adapter 60N-C 2/3" 0.63x

Nagy érzékenységû mikroszkóp kamera AxioCam MRm Rev. 3 FireWire (D) Mid Range Monochrome with AxioVision AC, FireWire / IEEE1394 interface incl. cable and IR barrier filter BG40 (enclosed) Number of Pixels: 1388 (H) x 1040 (V) = 1.4 Mega pixel Pixel size: 6.45 μm x 6.45 μm Chip size: 8.7 mm x 6.9 mm equivalent to 2/3" Spectral range: With protection glass app. 350 to 1000 nm With IR barrier filter app. 350 nm to 700 nm NIR-Mode: Increase of IR sensitivity Max. Full Well Capacity: Approx. 17.000 e S
electable Resolution: 
H x V Binning Factor

276 x 208   5 x 5

344 x 260   4 x 4

460 x 344   3 x 3

692 x 520   2 x 2

1388 x 1040 1 x 1

Live frame rates (depending on hardware and software configuration): 
H x V Binning Factor Frame Rate@20ms

1388 x 1040    1                 11

694 x 520        2                 21

462 x 346        3                 31

Harddisc recording: Inline recording of image data directly to harddisk at all speeds with AxioVision 4 Module "Fast Acquisition" available

Frame rates for time series recording in AxioVision 4 Module

"Fast Acquisition" (depending on special hardware requirements)

H x V Binning Factor Frame Rate

1388 x 1040 1                 14

692 x 520     2                 26

460 x 344     3                 35

344 x 260     4                 43

276 x 208     5                 50

Readout of Sensor Sub-Regions ("ROI"): Adjustable

Digitization: 12 Bit / 24.57 MHz pixel clock

Dynamic Range: Typical >2200 : 1 (>66.8 dB) at < 7.7 e readout noise

Dark current: ca. 0.7 e/p/s

Integration Time: 1 ms to 60 s

Cooling: One stage Peltier cooling

Control signals: TTL output for controlling of external shutters

Interface: FireWire / IEEE1394 interface, 6 pin jack, cable 5 m, 400MBit/s

Optical Interface: C-Mount

Thread depth for objectives: max. 5 mm

Max. file size per image: About 2.8 MB at 1388 x 1040 @ 1 x 12 Bit

Operating Systems: Windows 2000 Professional, Windows XP Professional

Size / Weight: App. 11 cm x 8 cm x 4.5 cm (2.3" x 3.2" x 1.7") / 370 g

Housing: Blue anodized aluminum, with cooling fins, 1/4" photo thread for tripod mount

Registration: CE, cUL

Power Supply: 10-33 V, DC, 4W, power supply provided by FireWire bus

Environmental conditions: 5° ... +35° Celsius,

10% .... 80% relative air humidity, not condensing, free air circulation required

Camera fényút kicsatoló oldalsó port, mot.,bal oldali, interface 60Nmotoros csúszkával 2 kapcsolási pozícióval, pos. 1 = 100% vis, pos. 2 az opcionálisan választható tükör felé

Opcionálisan választható tükör: 100% kamerához

AxioVision 4 Module Multi Channel Fluorescence. Module for the acquisition of different fluorescent dyes in independent channels

- Acquisition of up to 8 independent channels. Each channel can be acquired using independent exposure time and microscope settings. Image display can be toggled between single channel display or pseudo colored merge display. Channel colors can be freely chosen.
	1 db


